## Slide 1
Mechanical Measurements
Lesson 2

## Slide 2
The property A is called: «measurand», the reference property U is the «measurement unit»
the ratio between A and U today is always performed by a calibrated instrument !
the measurement a expresses the amplitude or intensity of the property A we are studying
Measurements with calibrated instruments: by far the majority of the measurements performed today. The instrument has the unit U memorized inside …
Direct measurements: when there is a direct comparison between the measurand A and the unit U
Indirect measurements: when a physical law is applied to measure quantities other than the one we are interested in:
i.e. measuring a velocity v = x/t by measuring a distance x and time t

## Slide 3
Direct instruments : (they do not transform the physical quantity - only very few)
Indirect instruments : they always transform the physical
quantity they acquire in input …
Analog instruments
Digital instruments
Deflection instruments
Null-out instruments
Recording instruments

## Slide 4
Basic MEASUREMENT CHAIN:

## Slide 5
First question to be answered: WHAT do we want to measure and WHY ?
It’s a “KNOWLEDGE PATH” that leads to new quantitative INFORMATION about the world surrounding us !
want we to have an approximate control of a physical quantity ?
want we to do a rigorous scientific measurement ?
want we to use the value of the measure for an automatic control ?
want we to know to what extent can a physical parameter vary to set an alarm signal ?
Quality of the measurement

## Slide 6
Length and displacement measurement
Many instruments are available to measure length and displacement in modern technology …
The most ancient is the graduated ruler:
Rulers have an approximation of 1 mm,
7.35
Vernier Calipers have an approximation of 0,05 mm (1/20 caliper) or 0.1 (1/10 caliper), over an average range of 200 mm !
Eg.

## Slide 7
Micrometer (Palmer): it is based on a “micrometric screw” with a pitch of 0.5 mm (sleeve). It has an approximation of 0,01 mm
There are 50 notches on the rotating screw (thimble), therefore for every notch the screw advances:
mm,
which is also the approximation of the instrument !
Of course the range is limited by the instrument screw opening …
Eg.

## Slide 8
we choose the instrument according to the “extension range” of the physical quantity to be measured, based on the amount of variation that is expected for the quantity and also on how fast the quantity changes its value during the measurement;
we read the numeric data on the output device of the instrument;
by means of the graduation of the instrument we associate the numeric data a with the U units, performing actually the real measurement: A = a × U
at this point we still only have a raw measurement of the physical quantity A. We have to identify the many uncertainties ε associated with the raw data, correct the data we obtained and then switch to the final measurement: A = (a ± ε) × U
with the diffusion of digital instrumentation is now possible to acquire many repeated measurements of a physical quantity a1 a2 ... an. This raises the question of identifying the true magnitude of the value. We have to go through a statistical analysis of the data …
Sensor or
Transducer
A   a
First question to be answered: WHAT do we want to measure and WHY ?

## Slide 9
We have to establish the QUALITY of a measurement
High quality measurements means being able to get data with low uncertainty; most of the times, this is also an expensive procedure !
To quantify the quality of a measurement we have to quantify first the quality of the instrument !
To this extent we can define 5 general metrological characteristics that apply to every measuring instrument or measuring system and completely describe its performances !
Measurement Span
Sensitivity
Stiffness
Accuracy
Measurement Rapidity (Dynamic Response)

## Slide 10
MEASUREMENT SPAN
The measuring span is the numerical range between the minimum and maximum values of the “measurand” the instrument can appropriately measure and within which the other four instrument characteristics are valid !
Span = high operation limit – low operation limit
Graduation curve :
The operation of every instrument is based on a physical law, this law is described by an equation; the same equation is also the equation representative of the graduation curve !
if this equation is of 1st degree, the instrument is said linear instrument !
if this equation is of 2nd degree, the instrument is said quadratic instrument !
…… overload

## Slide 11
Example: find the graduation curve of the Hg thermometer
the fundamental physical law (starting point) is the law of volumetric thermal expansion of all materials (fluids):
output input (measurand)
with
That’s the graduation curve !
Span ?
[-38 °C / +365 °C; …]

## Slide 12
Graduation curve should NOT be confused with the CALIBRATION (taratura) curve :
is the calibration curve: the curve of the differences between the reference value u and the raw instrument indication ub , for every instrument output ub .
if u - ub > 0 u > ub the instrument under calibration “underestimates” the measurand magnitude (input) so we have to add the indicated deviaton to ub .
if u - ub < 0 u < ub the instrument under calibration “overestimates” the measurand magnitude (input) so we have to subtract the indicated deviation from ub .
u is the reference value or the indication of a reference instrument
ub is the “raw” indication of the instrument under calibration

## Slide 13
SENSITIVITY
Capability of the instrument to «sense» small variations of the input variable (the measurand)
how small can a variation Δi of the input be, to get from the instrument an appreciable output Δu ?
We can certainly write that, for variations means , or better
It is immediately observed that, if u = f (i) is the graduation curve, the sensitivity is the derivative of the graduation curve !
The sensitivity S can be calculated for every point of the graduation curve, by means of the differential ratio:

## Slide 14
Examples:
Instruments with a constant sensitivity are called linear instruments !
Instruments with a sensitivity that is a function of i can be quadratic instruments !
Instruments with a sensitivity that is a function of i can be logarithmic instruments !
Do not confuse the sensitivity with the resolution of an instrument, which actually is a more appropriate concept for digital instruments !
geometric meaning: it’s the slope of the tangent to the graduation curve in the measuring point that is being considered.

## Slide 15
TH2O is the same in the bathtub and in the glass.
The loading effect of the instrument is bigger when measuring T in the glass than in the bathtub !
Example:
STIFFNESS
The introduction of any measuring instrument into a measured medium always results in the extraction of some energy from the medium, thereby changing the value of the measured quantity from its undisturbed state !
This circumstance makes a perfect measurement theoretically impossible !
Every instrument that physically interacts with the measured quantity has a «loading effect» on the measurand and slightly changes its value
Instrument designer can only «minimize» this loading effect, which is called stiffness !
Stiffness is a “singular characteristic” because it depends also on the measurement environment and circumstances.
An instrument with a low stiffness means an instrument capable of doing measurements with a small loading effect !

## Slide 16
Let’s go to a generalized definition of stiffness and input impedance :
Energy transfers in a measuring system require the specification of two variable quantities for their description.
At the input of each component there exists a variable qi1, the “effort variable”, that carries the “information content” of the measurement.
At the same input, however, there is a second variable qi2, the "flow variable" associated with qi1, in a way that the product qi1 × qi2 has the “dimensions of power” and represents an instantaneous rate of “energy withdrawal” from the preceding element !
Now the generalized input impedance can then be defined as :
The power drained from the preceding element is :
and a “large” generalized input impedance is needed to keep the power drain small !

## Slide 17
These concepts can be immediately applied to an electrical example: a voltmeter measuring an unknown voltage V.
As soon as the meter is applied to the voltage terminals, the electrical circuit is changed and the value of V is no longer the same but changes to another value Vm .
For a voltmeter the input variable of interest, the effort variable (qi1), is the terminal voltage Vm . If we look for an associated variable (qi2) which, when multiplied by qi1 , gives the power withdrawal from the voltage terminals, we find the meter current im meets this requirement.
Thus, the input impedance in this example is: the meter resistance !
This situation is very important indeed and applies also to the internal stages of a more complex instrument chain, for example, when coupling a transducer that measures a physical quantity A and outputs a voltage V or a current i to the signal processing stage that follows in the measurement chain ….

## Slide 18
Some numerical means of expressing the loading effect of the instrument on the measured medium would be helpful in comparing competing instruments at the moment of their choice or purchase !
One parameter could be the insertion or connection error done by the instrument when connecting with the measurand:
where: ab is the numerical value of the measurand “before” the instrument insertion;
a is the numerical value of the measurand “after” the instrument insertion, the actual measurement returned by the instrument.
Because ab is a value that can NOT be measured, this parameter may seem useless !

## Slide 19
Before connecting the voltmeter to the terminals we have :
V = V0 with I = 0
After connecting the voltmeter we have:
V0 = (Ri + Rv)·I with I ≠ 0
But the instrument measures:
V = Rv·I ≠ V0
If we wish to keep this error small we have to design either or
Acting on the transducer output resistance Ri in not so easy therefore, the preferred way in designing the connection is making the input resistance Rv of the signal processing stage (voltmeter) very big !
This choice prevents the signal V, which carries the “information” about the measurement, to degrade further !
Example I
We do a “connection error”

## Slide 20
Before connecting the ammeter to the terminals we have :
I0 = Ii with I = 0
After connecting the ammeter we have:
I0 = Ii + I with
And the instrument measures:
I ≠ I0
And the connection error becomes :
Example II
We do a “connection error”
Since the two resistors are now in a parallel configuration:
Therefore :

## Slide 21
If we wish to keep this error small we have to design either or
Having at disposal a current generator with a big internal short-circuit resistance Ri is advisable but not so easy to obtain from a transducer which outputs a current therefore, the preferred way in designing the two-stage connection is making the input resistance Ra of the signal processing stage (ammeter) very small !
This choice prevents the signal I, that carries the “information” on the measurement, to degrade further !