## Slide 1
Mechanical Measurements
Lesson 3

## Slide 2
ACCURACY (Precision)
Engineers and technicians always wished the instruments being able to perform precise measurements
(with high accuracy).
However it is unavoidable to make errors during measurement processes, so we must define a convenient way to express how good or how bad our measurements a are !
A first simple way could be expressing the “distance” between the measurement a and the true value av
ε = |a - av| which is the absolute error of the measurement.
However, there is a more convenient way in engineering to express the accuracy of a measurement:
Attitude of a measuring instrument to give the «true value» av of a physical quantity !
which is the relative error of the measurement (expressed in %), and gives an immediate and intuitive idea of the measurement accuracy !
If we knew the true value av , we were basically all set and we could go ahead to analyze the next characteristics.
Unfortunately, the true value of a physical quantity is generally unknown and unknowable !

## Slide 3
To get EXACT INFORMATION about the measurand we should have available an exemplar measurement method !
A measurement is therefore a technical procedure that strives to get as close as possible to the truth of the natural world and/or of the technology ! … But, it never succeeds …
We can only minimize the unavoidable uncertainties εa : trying to get εa → 0
There are two main types of error that can be done during a measurement:
Systematic (BIAS) errors
Random (PRECISION) errors
Measurement accuracy results from a combination of these two errors !
Bias errors are often “macroscopic” errors because they have a «technical origin» (instrument malfunction, bad conditions of use, strong external noise that the instrument can not reject, instrument out of calibration).
Bias errors CAN be discovered by a calibration procedure and, once highlighted, they MUST be eliminated !

## Slide 4
Random errors are generally small “scatter errors”, that can not be avoided during a measurement and for which is impossible to find an exact reason !
Random errors can be studied after the bias error has been corrected for, or (better) eliminated.
Therefore the sum of all random errors is, sometimes, also called residual error.
The residual error always depends on unknown small random reasons, therefore it is much better referred with the word uncertainty !
Accordingly to the way we evaluate it, there are only two different types of UNCERTAINTY, which are classified by the international «Guide to the expression of uncertainty in measurement» ed. JCGM 100: 2008 (GUM)
http://www.bipm.org/en/publications/guides/gum.html
Type A uncertainties: evaluated by a statistical analysis of series of observations (measurements)
Type B uncertainties: evaluated by a pool of comparatively reliable information

## Slide 5
The uncertainty of the result of a measurement reflects the “lack of exact knowledge” of the measurand value.
The result of a measurement, after correction for recognized systematic effects, is still only an estimate of the value of the measurand because of the uncertainty arising from random effects and from imperfect correction of the result for systematic effects.
In practice, there are many possible sources of uncertainty in a measurement, including:
incomplete definition of the measurand;
imperfect realization of the definition of the measurand;
non-representative sampling - the sample measured may not represent the defined measurand;
inadequate knowledge of the effects of environmental conditions on the measurement or imperfect measurement of environmental conditions;
personal bias in reading analogue instruments;

## Slide 6
finite instrument resolution or discrimination threshold;
inexact values of measurement standards and reference materials;
inexact values of constants and other parameters obtained from external sources and used in the data-reduction algorithm;
approximations and assumptions incorporated in the measurement method and procedure;
variations in repeated observations of the measurand under apparently identical conditions.
These sources are not necessarily independent: some of sources a) to i) may contribute to source j).
Of course, an unrecognized systematic effect cannot be taken into account in the evaluation of the uncertainty of the result of a measurement but contributes to its error !

## Slide 7
For simplicity, we start the study of uncertainty estimation with the type B uncertainty:
For an estimate xi of an input quantity Xi that has not been obtained from repeated observations, the
associated estimated variance u2(xi) or the standard uncertainty u(xi) is evaluated by scientific judgement
based on all of the available information on the possible variability of Xi .
The pool of information may include :
previous measurement data;
experience with or general knowledge of the behaviour and properties of relevant materials and instruments;
manufacturer's specifications;
data provided in calibration and other certificates;
uncertainties assigned to reference data taken from handbooks.
Examples of typical errors in mechanical measurements that can contribute to type B uncertainties :
ep = d tgφ
Reading error for analogic instruments

## Slide 8
Mobility error for mechanical instruments
Hysteresis error for mechanical instruments
Fidelity error: due to the disturbing and noise factors as temperature, humidity, electromagnetic fields, mechanical vibrations, atmospheric pressure, non-inertial reference frame ...
It quantifies the sensitivity of the instrument to external interferences, and it has to be assessed using multiple repeated measurements over time, by changing the position and the conditions of use of the instrument while keeping the input variable strictly constant …
Zero shift error: due to loss of calibration of the mechanical components (recording springs) or electrical components (trimmer), or to the aging of the electronic components (resistors and capacitors) …

## Slide 9
Calibration and standard references errors: due to the errors that are done while carrying out a calibration procedure or due to the inherent uncertainties of the reference standards used for the calibration !
±α are the standard references errors
±β are the errors done while tracing the calibration curve
The total error is a quadrature sum: and can be represented with the ellipsoid show in the figure.
An older way of expressing the type B uncertainties [in %] was with the precision class parameter:
where εi is each error or uncertainty source one is able to isolate and estimate
Example:
a class 0.5 dynamometer with a full scale of 100 N makes an absolute error of 0.5 N when measuring 100 N :
a relative error of: 0.5 N / 100 N = 0.5%
The same class 0.5 dynamometer makes the same absolute error of 0.5 N when measuring 5 N, which now is
a relative error of: 0.5N / 5N = 0.1 = 10 %.
An advice: never use an instrument for which a precision class is declared at the beginning of the span !

## Slide 10
We now proceed to explain how to estimate the type A uncertainty :
Type A uncertainty estimation has a completely different approach → it does NOT even care about identifying error sources !
The method is based on a statistical data post processing and, therefore, we need a sufficient number n of repeated measurements with the measurand X held strictly constant !
The final goal is NOT even to find the true value xv but a certain range of values within which the true value could reasonably be found ! This range could be written with where xb is the best representation of the true value (x best) and δx is a width parameter that might be representative of the uncertainty !
xi is the ith measurement of the measurand X
Because our n measurements x1, x2 … xn will be very similar each other but not equal between them, which of them could possibly be the best representation of our n measurements ?

## Slide 11
The mean value of our data: the n measurements !
So far, there is NO theoretical justification for our choice; its only a very reasonable choice, so reasonable that I will try to apply it also to find a quantitative expression for the width parameter δx :
It’s the “deviation” of the generic measurement xi from the mean
Can we calculate the mean value for this deviation ?
NO ! because the deviations are “equidistant” from the mean value !
Therefore, the Standard Deviation (of the population) has been proposed !
or, with only a few measurement points, the Standard Deviation (of the sample) ! which produces σx = 0/0 instead of σx = 0 when n → 1

## Slide 12
Now we can express our measurement xi with the best representative and with the width parameter σx !
We have now to ask ourselves:
How much do we trust these choices ? Is every measurement falling “inside” the range we just determined ?
Or, better, what CONFIDENCE do we place in this range of values ?
To answer these question we have first to construct the figure on the left, where we divided the “range” where the measurements fell in k small intervals Δk with at least one measurement inside each interval Δk .
The more measurements xi we have available the more (and smaller) intervals we can choose.
The height fk of each rectangle is proportional to the number of measurements inside each interval Δk .
This figure is called the Frequency Histogram.

## Slide 13
FREQUENCY HISTOGRAM
nk → number of measurements (observations) that fall into the interval Δk
→ frequency of the observations (measurements) that fall into the Δk interval
Obviously: and
The frequency histogram is “normalized” by definition !
The area represent the measurement fraction that falls within the Δk interval.
Note that it is possible to express the mean value of all the n measurements as:
where xk is the best representation of xi inside the Δk interval
We now think ideally to bring the number of measurements n → ꝏ … in this situation we can think of an “enormous” number of intervals Δk which are also going to be “smaller and smaller” …
The stepwise curve of the frequency histogram becomes a smooth limit distribution curve f(x)

## Slide 14
Where:
n → ꝏ
Δk → dx
fk → f(x)
f(x)dx is the measurement fraction that falls within the infinitesimal dx interval
is the measurement fraction that falls within the finite interval (b-a)
Because it is
the limit distribution curve is also normalized !
Only for n → ꝏ
Frequency
(experimental results)
Probability
(theoretical model )
But, now, which curve describes at best the limit distribution curve f(x) ?
It’s the well-known Normal distribution curve or Gaussian curve or bell curve …

## Slide 15
where X is the “mean” or the true value of x (in fact we have now n → ꝏ measurements)
and σ is a “width parameter” of the function
Applying the normalizing condition: we get
If we calculate the mean value for the function fX,σ (x) : we got the mean !
If we calculate the mean squared difference for fX,σ (x) : we got the variance !
which is the squared standard deviation …
Therefore, “theoretically” the mean X is the true value of the measurement and the width σ is the standard deviation, calculated for the ideal case of n → ꝏ measurements !!!

## Slide 16
Finally, if we refer to the mathematical model fX,σ (x), we can express the measurement x with :
We still have to understand what confidence we give to the width parameter σ …
The good news is that with a “mathematical model” we can calculate the integrals between two specified limits:
And they will represent the probability one measurement x has (one of the ꝏ we have at disposal) to fall within the range ± σ, or ± 2σ, or ± 3σ around the true value X
The factor k is named in the GUM as coverage factor and leads to the definition of expanded uncertainty …

## Slide 17
Therefore, writing means there is a 68.3 % “probability” to find a measurement xi within the uncertainty range and .
In other words, we gained a confidence of 68.3% of finding our measurements within the range .
The geometrical representation of the probability discussed above, is shown left on the gaussian distribution curve, also for coverage factors of 2 and 3, and is the area underlying the curve.
It is very important to understand the two following concepts:
is a probability and can be calculated before doing a single measurement (having the function) !
is a statistics and can be calculated only if we have data available (after doing the measurements) !

## Slide 18
Keep always in mind that the confidence intervals and their probability values have been determined by means of a theoretical model or considering ꝏ measurements available, which NEVER happens in the real world, where we ALWAYS have only a limited number n of measurements available !
Before applying the results of the probability theory, we have to be reasonably sure that the distribution of our measurement is actually a Gaussian one (χ2 test) otherwise, we should rather try applying other statistical distributions (t Student) …
Because, by increasing the number n of the measurements the standard deviation σx does not change much, is a good way of expressing the uncertainty of the instrument or of the measurement method.
By observing the shape of the measurement distributions, it is therefore immediate establishing or comparing the accuracy of different measuring instruments !!
Being we still have to understand if it is possible to get from our n measurements x1 x2 … xn any information about the accuracy of the measurement ?

## Slide 19
Distribution curve of the m mean values
Distribution curve of the n measurements
Standard Deviation expresses
the instrument accuracy !
Standard Error expresses
the measurement accuracy !
For the same set on n measurements, the Standard Error is always “smaller” than the Standard Deviation !!
It is the Standard Error of the mean and gives the uncertainty with which the mean represents the “true value” !
the Standard Error can be calculated using only the n measurements of one single group but, because of the way it was defined, it keeps the powerful meaning of UNCERTAINTY of the MEASUREMENT !!
if we make more measurements (n → big) the Standard Error decreases while the Standard Deviation stays about the same ! … making more acquisitions makes a better measurement, NOT a better instrument !

## Slide 20
Example: Discuss the accuracy / precision of the instruments A, B, C
6 measurements of a pipe diameter – nominal diameter 5.00 mm
A
5.01 mm
4.98 mm
5.00 mm
5.05 mm
4.97 mm
4.99 mm
C
5.03 mm
5.04 mm
5.03 mm
5.04 mm
5.05 mm
5.06 mm
B
5.10 mm
5.12 mm
4.95 mm
4.92 mm
4.94 mm
4.97 mm
0.02582 0.03 | 0.079373 0.08 | 0.010677 0.01
0.028284 0.03 | 0.086948 0.09 | 0.011696 0.01
Which are the measurement units?
5.00 5.00 5.042
5.04
Errore sistematico 0 0 0.04
CV=
A:0.6% B=1.8% C=0.2%

## Slide 21
Error Propagation Theory …
Consider a physical quantity “q = x + y” that can be expressed by the sum of two other primary physical quantities: and We wish to find:
with
which might represent the upper limit of the width δq
which might represent the lower limit of the width δq
We might, therefore, think that however, if the errors that lead to δx and δy are independent, the choice is an overestimation of the uncertainty for q …

## Slide 22
In fact, to actually have we should always underestimate or overestimate at the same time the measurement of x and y , which would imply an “underling low” or a “correlation” between the two variables, making them somehow dependent on each other !
If this is the case, then can be a correct choice otherwise, when the measurements and the errors of x and y are independent, there is always a “partial mutual deletion” of the uncertainties, the algebraic sum of the x and y errors is then an “overestimation” of the q uncertainties and it is much more reasonable to add the errors for x and y in quadrature :
Similarly, when we have a measurement expressed as a product or a quotient of two primary quantities the general rule is to add the relative errors in quadrature :

## Slide 23
In general, when the variable of interest q is a function of a measurable physical quantity x : q = q(x)
(for ex. )
It is always possible to measure and calculate with the functional relationship !
But how are we going to calculate δq ?
If δx is small and due only to random errors, qmin and qmax are almost equidistant from qb of a distance δq, regardless of the function type !
In this hypothesis we can write:
which for δx → small, can also be written as:
which is the derivative of the function q(x) calculated in xb
We reach then the important relationship:

## Slide 24
To include also the common cases when the function q(x) is decreasing in xb and has a negative derivative in the point xb :
we should rather adjust the result and consider the absolute value of the derivative …
In general, when we wish to know the intensity of a physical quantity “q” that can be expressed with a function of two or more other physical variables q = q(x,y) , and we measure these primary variables with their uncertainties:
and ; then we can always use the function relationship to calculate the best representative of q :
while for δq we might consider to apply the superimposition of the effects and use the algebraic sum:

## Slide 25
However, again, if the measurements of x and y and their errors are independent, it is quite reasonable to consider that there will be a “partial mutual deletion” of the uncertainties and, to express the general uncertainty δq it is much more reliable to add the uncertainties of x and y in quadrature:
Please, note that the width parameters we used for the measurements x and y are nothing more than the Standard Deviations previously calculated: and
Everything said so far can be generalized for the case of a physical quantity q = q( x1 , x2 , … xn ) which is a function of n measurable primary independent variables:
Which is the fundamental statistical law of the Uncertainty Propagation and represents also the Combined Standard Uncertainty for uncorrelated input quantities reported in the GUM – JCGM 100: 2008

## Slide 26
We are now ready to answer our last question:
Can we actually estimate somehow the accuracy of the measurement itself with only a “limited number N” of measurements x1 x2 … xN ??
This goal can be approached by estimating “how well” the mean value represents the true value X or, trying to calculate the uncertainty of when the mean value is the best representative of the true value X .
To do so, we start dividing our N measurements xi in m groups of n measurement each :
Now we have N = n × m measurements and, for each group of n measurements, we can calculate the mean value (j = 1, 2 … m)
Who is now the best representative of the m mean values ??
It’s the mean value of the m mean values :
If X is the “true value” of the n × m measurements , it is also the “true value” for the m means because
they come from the same measurements !
the “mean of the means”

## Slide 27
If the n × m measurements are affected only by small random errors, each distribution curve for the m groups of n measurements will be a normal (Gaussian) distribution curve pj (x) .
Therefore, the m mean values will also distribute with a normal (Gaussain) curve around the “mean of the means”.
This happens because each mean value is a function of the n measurements :
The width parameter for the distribution of the m mean values will be :
the Standard Deviation of the mean
The width parameters δxj are here the Standard Deviations calculated with the measurements of the jth measurements group. Since are all measurements of the same quantity done with the same instrument …

## Slide 28
… for a “reasonable number n” of measurements, they will be almost coincident between them :
From the definition of mean value: it is easily obtained:
And substituting all these positions in the main equation of the width parameter it results:
It is the Standard Error of the mean and gives the uncertainty with which the mean represents the “true value” !
Please, carefully note:
the Standard Error can be calculated using only the n measurements of one single group but, because of the way it was defined, it keeps the powerful meaning of UNCERTAINTY of the MEASUREMENT !!
if we make more measurements (n → big) the Standard Error decreases while the Standard Deviation stays about the same ! … making more acquisitions makes a better measurement, NOT a better instrument !