## Slide 1
Mechanical Measurements
Lesson 4

## Slide 2
RAPIDITY (Dynamic Response)
So far the measurand (the physical quantity for which we wish to measure the intensity) has been considered strictly CONSTANT during the whole measurement procedure !
From here on it will change its intensity with time during the measurement … the measurand becomes a function of time !
i(t) Instrument u(t)
both the input measurand and the output measurement are functions of time → i(t) u(t)
We may call rapidity of an instrument “the attitude to correctly follow the changes of the measurand during time”
Rapidity of mechanical instruments is always limited by the inertia and the damping effects of its moving parts !
Rapidity for electronic instruments is always limited by the combination of its capacitive and inductive reactances !
An instrument with insufficient rapidity (insufficient dynamic response) during a measurement will output a measurement waveform which will be attenuated and out of phase (delayed) with respect to the measurand !
The output measurement waveform will be distorted with respect to the input measurand waveform …

## Slide 3
We have basically two schemes or methods to study the dynamic response of a measurement instrument:
STEP RESPONSE with which we can get some important parameters such the settling time and the time constant
FREQUENCY RESPONSE with which we can get other important parameters such the cut-off frequency and the damping factor …
There is also a third scheme, which is used less often: the ramp response that leads to the delay time …

## Slide 4
At time t0 the input signal changes instantaneously its value from i0 to i1
The instrument output will try to follow this change, form u0 to u1 , the best he can…
However, it will take some time ts = t1 – t0 to reach the correct final value within a certain error ±εd !
±εd is the dynamic error, which has to be established in advance by the operator …
ts = t1 – t0 is the settling time of the instrument, which provides a first idea of its dynamic response
tSLEW = tsr – t0 is the slew rate, which is the time the instrument takes to reach the first overshoot peak…
… not all the instruments show an overshoot during the step response !
Step Response

## Slide 5
In the more general case, an input variable (measurand) changes “periodically” during time, then we can refer to simple sinusoidal inputs because every periodic signal can be decomposed by the Fourier Series:
input: with frequency
output: with phase delay
ideal frequency response
and phase response curves:
constant output/input ratio for “every frequency” of the input !
zero phase delay for “every frequency” of the input !
Frequency Response

## Slide 6
Real instrument frequency response example:
Real instruments always attenuate the output amplitude for higher frequencies (f → ꝏ) because:
mechanical instrument have inner moving parts with inertia which can not have “infinite acceleration”
electronic instrument have inner components which can not have
inductive reactance XL = jωL = ꝏ
capacitive reactance XC = 1/jωC = 0
The extension of an instrument frequency response depends also from the dynamic error the operator accepts:
if we accept a dynamic error εd = 3% we have a cut-off frequency f2
If we accept a dynamic error εd = 25% we have a much higher cut-off frequency f1 > f2
All frequencies between f0 and f1 (or f2) form the instrument bandwidth or pass-band !

## Slide 7
There are many instruments where the bandwidth does “not start” from frequency f0 = 0 Hz, but starts at a higher value f1
In these cases, we have “two cut-off frequencies” and the pass-band is between these two frequencies:
Most of the times the dynamic error (which can be “fairly large”) is expressed in a logarithmic scale, inherited from “acustics” studies and the logarithmic unit is the decibel :
It is world wide accepted that a tolerable dynamic error limit is which means the instrument output A is “attenuating the input” to the 70,7 % of its real amplitude or that it is doing a dynamic error of 29,3 % !!

## Slide 8
We say an instrument is dynamically linear if we can write its “dynamic characteristic equation” as follows :
The instrument dynamic can be described by an ordinary differential equation with constant coefficients !
The solution can be written as:
Solution of the associated homogeneous equation which represents: xtr(t)
the transient dynamic response of the instrument
Solution of a particular integral of the equation which represents: xrg(t)
the dynamic steady state of the instrument
y(t) Instrument x(t)
We change now notation for the input and output variables:
y(t) → input variable (measurand)
x(t) → output variable (instrument response or
indicator deflection)

## Slide 9
To study the TWO different situations we will employ the TWO different experimental schemes or tools introduced before:
For the dying out transient → the STEP RESPONSE xtr(t)
For the established steady state → the FREQUENCY RESPONSE xrg(t)

## Slide 10
Zero-order instruments
When the differential equation of the dynamic characteristic has the simplest form, without the derivative terms:
→
The output signal (or deflection) is “always” proportional to the input signal !
There is no delay between output and input therefore, zero-order instruments “respond instantaneously” and can be considered ideal instrument !
For the potentiometer (as a position transducer) we have for the left circuit E = RI , while the displacement y is indicated by the output voltage e(t) which is:
where and the variable “time” is only
present as “independent variable” for the position y = y(t)
Note that is also the graduation curve (static characteristic
curve) of the potentiometer and is the static sensitivity !
Example:
e(t)

## Slide 11
First-order instruments
A first-order instrument can “store energy” in only one of its inner elements !
If we displace the indicator S of the simple mechanical system to a position x0 ≠ 0 , elastic (potential) energy is stored in the spring k.
If at t = 0 we leave the indicator S free of moving, the elastic energy stored in the spring will pull back the indicator to the rest position
x = 0 . During movement, the damper c will oppose the motion.
For every time instant t the equation of force equilibrium is: or
For every time instant t the velocity is: therefore the coefficient is a time !
is the time constant of the first-order instrument of above !

## Slide 12
The solution of the 1st order differential equation is the well known “decreasing exponential”:
the curve of which can be drawn with the initial condition: t = 0 →
- The tangent line to the curve in the point t = 0 is:
Observe in the figure the geometrical meaning of the time constant
- At the time t = λ we have that the indicator S has travelled for 63% of the displacement and has still a 37% to go to reach its final position x = 0 !!
… but these are only introductory definitions and considerations …

## Slide 13
In the general case of the step response, an input variable is present : a step force F0 is applied to the indicator S at time t = 0 …
The 1st order differential equation that describes the indicator deflection therefore is: and the general solution has the form
xrg is a particular solution and describes the steady state condition of the instrument:
Note that when t → ꝏ the indicator speed must be zero , therefore the steady state of the
indicator S will be described by the position
xtr is the solution of the associated homogeneous equation and describes the dying out transient:
Observe again on the figure the meaning of
At t = λ we have
The final solution will be the “sum” of the two contributions:
F, x

## Slide 14
When we are interests in studying the steady state dynamic response of a 1st order instrument we employ the other tool, the frequency response. To do so, we have to apply at the instrument input a periodic measurand: a periodic force of frequency ω = 2πf !
The differential equation which describes the indicator S movements will be:
The general solution will always be but we will be interested only in the steady state part, which we can assume to be:
How does the “amplitude” of the output indication x(t) change with frequency ? And what “delay” will the output x(t) show ?
Let's shift to the phasor notation:
IN:
OUT:

## Slide 15
To verify the xrg solution we have to substitute our assumed “steady state solution” in the differential equation …
where is the maximum amplitude and is the time constant !
… calculating the modulus of the rationalized function, we will obtain:

## Slide 16
which is the amplification or gain of the instrument !
which is the phase delay of the instrument !
Output amplitude / maximum amplitude

## Slide 17
1st order instrument frequency response
1st order instrument phase delay
These diagram have on the x axis the non-dimensional variable “ωλ» and therefore are normalized diagrams
is the characteristic frequency of a 1st order instrument for which we have ωcλ = 1 →
therefore, ωc is the cut-off frequency with a - 3db attenuation :
 0
 
=1/

## Slide 18
Example:
The “dynamic response” of the mercury thermometer
For every differential time interval dt that precedes the equilibrium condition we have:
Heat amount that the environment “transfers to” the mercury:
Heat the mercury “receives from” the environment:
Where: m is the total mercury mass
c is the mercury specific heat
k is the thermal exchange coefficient
A is the thermal exchange surface of the thermometer
The two heat amounts of above must be the same:
Then the time constant for this 1st order instrument is
If λ = 1s then ωλ = 1 → and → quite slow !

## Slide 19
Second-order instruments
A second-order instrument can “store energy” in two of its inner elements !
k → the elastic spring stores potential energy
m → the moving mass stores kinetic energy
The differential equation that describes the “free movement” of the indicator m
If we consider the very particular case for which c = 0 we get the important limit situation described by :
where and ωn is the natural frequency of the instrument !
A system like this has NO damping and, if pulled from its equilibrium position x = 0, it will oscillate indefinitely with a frequency ωn and will NEVER stop !
Of course, it’s an IDEAL system …

## Slide 20
In the real case of the step response, an input variable is present : a step force F0 is applied to the indicator m at time t = 0 …
The 2nd order differential equation that describes the indicator deflection therefore is: and the general solution has again the form
xrg is a particular solution and describes the steady state condition of the instrument:
Note that when t → ꝏ the indicator speed and acceleration must be zero therefore the steady state of the indicator S will be described by the position
xtr is the solution of the associated homogeneous equation and describes the dying out transient:
with and the discriminant !
Depending on the sign of the discriminant Δ > 0, Δ = 0, Δ < 0 we have “different behaviors” of the instrument
The 2nd order instrument will exhibit “more than one type” of step response !
The curves of the step response will depend on a parameter obtained from the discriminant Δ :
the damping factor where

## Slide 21
For ξ > 1 (which is the same of c > ccr ) the instrument is over-damped and it will reach the final indication with a slow exponential path !
For ξ = 1 (which is c = ccr ) the instrument has a critical damping and reaches the final indication with the fastest exponential path !
For ξ < 1 (which is c < ccr ) the instrument is under-damped, it reacts quickly to the step input but exhibits several damped over-oscillations before reaching the final indication !
Note that for the case ξ < 1 the frequency of the over-oscillations seems to depend from the value of the damping factor ξ …
Instrument designer generally prefer to choose a damping factor ξ = 0.7÷0.8 so the indicator m will have only one over-oscillation and reaches the final indication F0 /k from the upper side …

## Slide 22
To obtain the values of ωn and ξ one should know the values of the coefficients m, c, and k of the 2nd order instrument which is often almost impossible !
In case the instrument has ξ < 1 (which is the most common) one can employ an experimental procedure that leads to the logarithmic decay δ :
Displace the instrument indicator from the zero position, leave it free to move back and measure the decreasing oscillations …
If the decay is small, it can also be measured “after k waves”, as showed in the equation above.
The logarithmic decay δ is dependent on the damping factor ξ :
And the frequency ω0 = 2π/T0 of the decreasing oscillation is “smaller” than the natural frequency ωn … depending also from the damping factor ξ :

## Slide 23
When we are interested in studying the steady state dynamic response of a 2nd order instrument we have to employ the other tool, the frequency response. To do so, we have to apply at the instrument input a periodic measurand: a periodic force of frequency ω = 2πf !
The differential equation which describes the indicator m movements will be:
The general solution will always be but we will be interested only in the steady state part, which we can assume to be:
How does the “amplitude” X0 of the output indication x(t) change with frequency ?
And what “phase delay” ϕ will the output x(t) show ?
As previously done for the 1st order instruments we substitute the input and the instrument output in the differential equation:

## Slide 24
with
which is the complex frequency response for the 2nd order instrument
Calculating the modulus and the phase lag of this function we obtain …

## Slide 25
The amplitude of the frequency response of the instrument:
These curves are also dependent from the damping factor ξ
For the amplification is:
wich for «small values» of ξ can be «dangerous», these are the resonance conditions …
ξ : 0.707
Normalized diagram of the amplification or gain of the instrument:

## Slide 26
the phase delay of the instrument output !
The curves depend on the damping factor ξ , however they all pass through the point :
which makes it useful to experimentally determine the natural frequency ωn without knowing the damping factor ξ …
If you consider the steady state output for the step response, the static sensitivity will result:
which makes the dynamic response of an instrument always inversely proportional to its static sensitivity !!

## Slide 27
Example: the galvanometer
Static characteristic (graduation curve)
motor torque:
resisting torque:
the sensitivity is constant and
inversely proportional to ωn
Dynamic characteristic (dynamic response)
The output indication is a rotation θ therefore : and for t > 0 (step response)
(frequency response)