## Slide 1
Mechanical Measurements
Lesson 8

## Slide 2
Analog to Digital (A/D) Conversion:
Electronic and Digital Counters are fundamental devices for Analog to Digital and Digital to Analog Conversion
Sampling a signal means to “extract”, from the analog continuous voltage signal, a certain number of values (the samples), measured at a regular time distance (the sampling period) !
Sampling a signal is always a loss of information !
The inverse of the sampling period is the sampling frequency :
If the sampling frequency is not “chosen carefully” the information of the analog measured signal can be completely spoiled !
The “sampling” procedure is done by special electronic devices called A/D Converters

## Slide 3
Sampling a signal …

## Slide 4
AO as a comparator:
if Vref is conneced to - , we have the non-inverting comparator (Vin)
if Vref is conneced to + , we have the inverting comparator (Vin)
VOUT = A (V+ – V-)
Vcc
-Vcc

## Slide 5
Today there are many different types of A/D converters employed in so many appliances of consumer electronics.
We will study the three most important A/D converters for measurement instruments :
The integrating converter or voltage to time converter is the device which equips all the bench digital voltmeters in the laboratory !
The analog signal vi must be kept constant for all the time of the conversion; this task is performed by an external device the sample and hold circuit :
1 integrator
1 comparator
1 reference voltage generator
a cirtcuit for time measurement

## Slide 6
The logic unit connects the constant input signal vi to the integrator for a fixed time T , the output voltage vo will be:
After the fixed time T , the logic unit switches the integrator input to an internal signal -vr , whose value has always the opposite sign of the input signal.
Both signals vi and -vr are constant and produce two ramps at the integrator output vo …
The integrator output voltage vo at time T reflects the charge accumulated by the capacitor C during that time, which is also the charge the same capacitor releases during the time tm of the descending ramp:

## Slide 7
Therefore: , the device characteristic curve !
Because vr and T are constant for every input signal vi , the signal vi is simply proportional to the measured time tm which, in fact, is different for each input voltage vi1 , vi2 , …
The time tm is measured by the digital counter, starting when the logic unit connects the internal reference signal and stopping when the “comparator signals to the logic unit” the integrator output vo has crossed zero !
High resolution (sensitivity) due to the electronic counter / Low frequency of acquisition (around 50 Hz) - The sampling period is at least : T + tm

## Slide 8
Successive approximation converter
3 bit
0112 = 3/8 kg
1/2kg=0; 1/4kg=1; 1/8kg=1 (left)
1012 = 5/8 kg
1/2kg=1; 1/4kg=0; 1/8kg=1 (right)
½ kg is too much
½ kg is too low
¼ kg is too low
¼ kg + ½ kg is too much
3/8 kg is OK
5/8 kg is OK

## Slide 9
4 bit example
This converter is based on the summing Operational Amplifier and the working principle is based on comparing the input signal vi with “successively added” internal reference voltages.
The logic unit operates successively on the switches bi with i = 1, 2, … 4, which can assume the value 1 (close) or 0 (open).
Successive approximation converter
1 summing circuit
1 comparator
a reference voltage
The characteristic curve of the device is therefore:

## Slide 10
The input unknown signal to the device is Vi : at the beginning all the switches are open
the characteristic curve is :
the logic unit closes b1 : it results the switch remains close: b1 = 1
the logic unit closes b2 : it results the switch will be reopened: b2 = 0
the logic unit closes b3 : it results the switch remains close: b3 = 1
the logic unit closes b4 : it results the switch remains close: b4 = 1
where Vk k = 1…4 is the summing OA output voltage at each step.
At the end of the “successive approximations” procedure we get :
and the corresponding binary code is “1 0 1 1”
Note that the last step ! There is a discretization error
The discretization error is due to the limited number of switches (bits) of the A/D
converter, in fact the resolution of the device is equal to the LSB :
The procedure goes step by step, as in the next example :

## Slide 11
Discretization error can be minimized only adding more bits (switches) to the A/D converter !
Modern A/D converter have at least 12 bits or 16 bits, which means 216 – 1 = 65.536 discretization levels !
Note also that successive approximation converters always approximate by defect the analog value, and to mitigate this problem the characteristic curve is shifted ½ LSB to the left, as shown below for a 3 bit converter :
The figure show also another unavoidable error the A/D converters do, the saturation error : when all the bits (switches) are set to 1 the approximated output value is 7/8 of the Full Scale voltage or range : VFS =2Vref .
This problem can be solved only by increasing the range of the A/D converter !
Successive approximation A/D converters have a fixed machine time to do the conversion therefore, the sampling period is dependent only on the number of successive approximations (bits) and is relatively short (a few μs ), which makes this A/D converter a relatively fast device !
1/2LSB=0.625V
7.5V-0.625V=6.875V  6.875V-7.3V =-0.425V
6.25 7.5
0 1/8 2/8 3/8 4/8 5/8 6/8 7/8
111
110
101
100
011
010
001
000
7.3 V
FS=10V
6.25V-7.3V=-1.05V

## Slide 12
The fastest A/D converter employed in measurement instrumentation is the FLASH converter :
The resistor network operates a “voltage partition” from the full range Vr to zero. The comparators put to 0 (low) the output if the input signal Vi is lower than the partitioned inner reference voltage, or to 1 (high) if the input signal Vi is higher than the corresponding partitioned voltage !
The comparators work in parallel and do the comparisons all together. The codifying network reconstructs the binary code for the corresponding input voltage level !
2n – 1 comparators
2n resistors
Flash comparators are incredibly fast (Tc < 1 μs) but are increasingly complicate when the bits increase: an 8 bit comparator requires 28 – 1 = 255 comparators inside.
u7
u6
u5
u4
u3
u2
u1

## Slide 13
Binary output
u7
u6
u5
u4
u3
u2
u1
Eg. VX = 0,45VR
Then VX > 1/16, 3/16, 5/16, 7/16 VR
But VX < 9/16 VR
 First 4 comparators stay high
The comparators give in output: 0001111.
This is a digital output but needs to be decoded into a "3-bit" binary output
9/16=0.5625
7/16=0.4375

## Slide 14: AD converter resolution
100
200
150
50
0
Time (ms)
0
1.25
5.00
2.50
3.75
6.25
7.50
8.75
10.00
Amplitude
(volts)
16-Bit Versus 3-Bit Resolution
(5kHz Sine Wave)
16-bit
3-bit
000
001
010
011
100
101
110
111
|
|
|
|
|
3 bit : 2^3 = 8
16 bit : 65536
0-10 V input range  ~150 uV

## Slide 15
The velocity of conversion is directly related to the dynamic response of such devices :
If the conversion procedure is fast, the sampling period Ts can be short and the sampling frequency fs can be high!
The Nyquist sampling theorem states that the sampling frequency of the A/D converter must be at least double of the highest frequency of the signal :
If this rule is not observed, the discrete signal will be affected by aliasing:
Be advised that, to have a good digital reconstruction of the analog waveform, it is good practice to use a sampling frequency at leest 10 times higher than the signal highest frequency !

## Slide 16
Original 10 Hz sine wave signal
fc=100 Hz
fc=12 Hz
fc=27 Hz

## Slide 17
Original 10 Hz sine wave signal
fc=100 Hz
fc=12 Hz
fc=27 Hz

## Slide 18: Example
100
200
150
50
0
Time (ms)
0
1.25
5.00
2.50
3.75
6.25
7.50
8.75
10.00
Amplitude
(volts)
Different Gains for 16-bit Resolution
(5kHz Sine Wave)
Gain = 2
|
|
|
|
|
Your Signal
Gain = 1
    Input signal = 0 - 5 Volts
    ADC Range = 0 - 10 Volts
    Gain = 2
Amplified Signal