## Slide 1: MEASUREMENTS FOR INDUSTRIAL SUSTAINABILITY AND ENVIRONMENTAL SAFETY
LESSON 2

## Slide 2: Exercise 1
A pressure gauge with a measurement range of 0-10 bar has a quoted accuracy of 1.0% of the full-scale reading.
What is the maximum measurement error expected for this instrument?
What is the likely measurement error expressed as a percentage of the output reading if this pressure gauge is measuring a pressure of 1 bar?
MEASUREMENTS FOR INDUSTRIAL SUSTAINABILITY AND ENVIRONMENTAL SAFETY

## Slide 3: Sol. 1
MEASUREMENTS FOR INDUSTRIAL SUSTAINABILITY AND ENVIRONMENTAL SAFETY
The maximum error expected in any measurement reading is 1.0% of the full-scale reading, which is 10 bar for this particular instrument. Hence, the maximum likely error is 1.0%x10 bar = 0.1 bar.
The maximum measurement error is a constant value related to the full-scale reading of the instrument, irrespective of the magnitude of the quantity that the instrument is actually measuring. In this case, as worked out above, the magnitude of the error is 0.1 bar. Thus, when measuring a pressure of 1 bar, the maximum possible error of 0.1 bar is 10% of the measurement value.

## Slide 4: Exercise 2
The width of a room is measured 10 times by an ultrasonic rule and the following measurements are obtained (units of meters):
5.381 5.379 5.378 5.382 5.380 5.383 5.379 5.377 5.380 5.381
The width of the same room is then measured by a calibrated steel tape that gives a reading of 5.374 m, which can be taken as the correct value for the width of the room.
What is the measurement precision of the ultrasonic rule?
What is the maximum measurement inaccuracy of the ultrasonic rule?
MEASUREMENTS FOR INDUSTRIAL SUSTAINABILITY AND ENVIRONMENTAL SAFETY

## Slide 5: Sol. 2
The mean (average) value of the 10 measurements made with the ultrasonic rule is 5.380 m. The standard deviation ±0.003 m. Thus, the precision of the ultrasonic rule can be expressed as 0.003 m (3 mm).
The correct value of the room width has been measured as 5.374 m by the calibrated steel rule. All ultrasonic rule measurements are above this, with the largest value being 5.383 m. This last measurement is the one that exhibits the largest measurement error. This maximum measurement error can be calculated as: 5.383 - 5.374 = 0.009 m (9 mm). Thus, the measurement accuracy can be expressed as +9 mm.
MEASUREMENTS FOR INDUSTRIAL SUSTAINABILITY AND ENVIRONMENTAL SAFETY

## Slide 6: Exercise 3
A packet of resistors bought in an electronics component shop gives the nominal resistance value as 1000 Ω and the manufacturing tolerance as ±5%. If one resistor is chosen at random from the packet, what is the minimum and maximum resistance value that this particular resistor is likely to have?
MEASUREMENTS FOR INDUSTRIAL SUSTAINABILITY AND ENVIRONMENTAL SAFETY

## Slide 7: Sol. 3
The minimum likely value is 1000 Ω – 5% = 1000 Ω – 50 Ω = 950 Ω.
The maximum likely value is 1000 Ω + 5% = 1000 Ω + 50 Ω = 1050 Ω
MEASUREMENTS FOR INDUSTRIAL SUSTAINABILITY AND ENVIRONMENTAL SAFETY

## Slide 8: Exercise 4
A particular micrometer is designed to measure dimensions between 50 and 75 mm.
What is its measurement range?
MEASUREMENTS FOR INDUSTRIAL SUSTAINABILITY AND ENVIRONMENTAL SAFETY

## Slide 9: Sol. 4
The measurement range is simply the difference between the maximum and minimum
MEASUREMENTS FOR INDUSTRIAL SUSTAINABILITY AND ENVIRONMENTAL SAFETY

## Slide 10: Exercise 5
Suppose that the instrument characteristic shown in figure is that of a pressure sensor, where the input units are expressed in bars from 0 to 9 bars and the output units are expressed in volts from 0 to 13 V.
What is the maximum nonlinearity expressed as a percentage of the full scale deflection?
What is the resolution of the sensor as determined by the instrument characteristic given?
How is the sensitivity?
MEASUREMENTS FOR INDUSTRIAL SUSTAINABILITY AND ENVIRONMENTAL SAFETY

## Slide 11: Sol. 5
The maximum nonlinearity is the maximum deviation of any data point on Figure is 0.5 units, which translates to 0.5 V. The full scale deflection (calculated for the fitted straight line) is 13.0 units, which translates to 13.0 V. The maximum nonlinearity can therefore be expressed as (0.5/13)%= 3:8% of the full scale deflection.
The resolution of the sensor as determined from the graph is the smallest change in input that is detectable. For the graph paper illustrated, is one tenth of a unit or 0.1 bar.
The Sensitivity is dO/dI= 13/9 V/bar
MEASUREMENTS FOR INDUSTRIAL SUSTAINABILITY AND ENVIRONMENTAL SAFETY

## Slide 12: Exercise 6
The following resistance values of a platinum resistance thermometer were measured at a range of temperatures.
Determine the measurement sensitivity of the instrument in Ω/℃
MEASUREMENTS FOR INDUSTRIAL SUSTAINABILITY AND ENVIRONMENTAL SAFETY
Resistance (Ω) | Temperature (℃)
307 | 200
314 | 230
321 | 260
328 | 290