## Slide 1
Mechanical Measurements
Lesson 10

## Slide 2
Measurement of Mechanical STRAIN
Strain measurements are perhaps the most widespread measurements done
in engineering (tension, force, pressure, …):
We measure here the lengthening and shortening of mechanical structures. If we know the Young’s Modulus E of the material, after measuring the strain εa we can obtain the value of mechanical tension σ = E⋅ ε !
Transverse strain is measured by considering the radial contraction and involves the knowledge of the
Poisson Coefficient:
Which for metallic materials ranges from 0,2 up to 0,4 …
Conventionally ν = 0,3
-

## Slide 3
Ductile and brittle materials:
A first type of instruments used to measure the strain or the displacement between two adjacent points are the
Mechanical Extensometers …

## Slide 4
Mechanical Extensometers :
The displacement is measured between two points, initially at distance l0
When the specimen elongates Δl, the indicator rotates of an angle θ and we read the output λ …
The ratio is :
And the graduation curve is then:
"not-punctual" strain measurement
no dynamic due to the masses
High amplification:
Static characterization / calibration


## Slide 5
However, what we really need is a punctual strain measurement, so to effectively monitor the zones of the surface where the strain concentrates.
Electrical strain gages are based on resistance variation with elongation !
Electrical STRAIN GAGES :
(historical realization)
(modern realization)
First we have to elaborate a bit the basic physical relation of resistivity …
Low sensitivity
Notes: Costantana

## Slide 6
then we derivate the relation and we transform the infinitesimal terms into finite differences …
Because it is and
it remains ( εt sign is the opposite of εa )
The sensitivity
with we would have
however, variation of Δρ/ρ is not zero !

## Slide 7
Commercial strain gages :
Because the contribution of the resistivity variation Δρ is difficult to calculate, strain gage producers proceed experimentally and «measure» the factor F for a certain number of gages.
The resulting number is then assigned to the same lot of transducer … that is the gage factor
Strain gage «graduation curve»
Note that Δρ/ρ is almost
constant for strains up to 5-6 % !!

## Slide 8
Thermal effects :
Unfortunately, because an electric current is flowing through every strain gage, there is an additional thermal
contribution on resistance variation:
Due to heating, the strain gage wire undergoes also a thermal elongation which, moreover, is
generally NOT equal to the thermal elongation of the underlying material We have therefore:
a resistance variation due to mechanical strain:
a resistance variation due only to thermal effects:
There is always an “apparent strain” which, in fact, is NOT real :
Because the coefficients α and β’ depend on the gage wire material while the coefficient β’’ depends on the underlying material, if we can make then we are nulling the apparent thermal contribution εapp
Temperature self-compensated strain gages approximately realize:
"constrain strain"
Notes: a: coeff termico di resistenza

## Slide 9
Temperature self-compensation is not perfect, nor constant with temperature, and it works ONLY if the strain gage is applied on the “correct underlying material” …
If how much is actually the resistance variation if we measure a strain of
100 μm/m with a 120Ω gage ? …
It’s a variation of only the 0,02% of the base resistance ! These small variations are optimal to be measured with the Wheatstone Bridge …

## Slide 10
We have now a combined measurement chain: “Strain Gage + Wheatstone Bridge”
The two graduation curves work then together …
Which combined give us:
the combined graduation curve !

## Slide 11
Strain gage (2) does not measure any strain but, it undergoes the same thermal effects of strain gage (1)
If the two strain gages (1) and (2) are applied on contiguous arms of the Wheatstone Bridge, the thermal effects
are “automatically eliminated”:
This is perhaps the most important service the WB does during strain measurement !

## Slide 12
The second gage (2) can be applied more efficiently on the specimen to increase the measurement sensitivity, but it must be rotated 90° in a XY configuration to give a useful signal. It will then measure the transversal strain producing the graduation curve:
where the factor (1 + ν) = 1,3 is an extra amplification due only to the bridge configuration, called «bridge factor»
Measurement of tensile strain:
The bridge factor can be increased up to 2(1+ν) if we employ four strain gages

## Slide 13
Measurement of bending strain:
the upper gage measures the stretched fiber strain while the bottom gage measures the compressed fiber strain :
because
The “bridge factor” is equal to 2 !
The bridge factor can be increased up to 4 if we employ four strain gages connected as shown on the right and make the full bridge active !

## Slide 14
Strains, of course, can be composed in more complicated ways , as shown in the figures ...

## Slide 15
Measurement of torsion strain:
Fundamental application for the crankshafts, where we wont to measure the drive torque to get the power output of an engine:
Strain gages are placed as in the figure, aligned with the maximum strain 45° lines on the shaft surface. The strains will be and the graduation curve results:
with a full bridge factor equal to 4 !

## Slide 16
Examples of complete strain gage measurement chains :
Note that strain gage channels need to be calibrated before use, this is done by connecting internally shunt resistors Rs that simulate specific strain values:
with 
Strain gage
Completion circuit
Power supply
gain
sensor
Matching circuit
output
 =
and