## Slide 1
Mechanical Measurements
A.Y. 2020-21
Lesson 13

## Slide 2
The measurement of the TEMPERATURE
If a certain amount of heat Q is supplied or subtracted from a body, this will change its temperature T.
The heat Q (which is a form of energy and is measured in Joule) alters the molecular activity of the bodies, causing variation of the temperature T. Therefore, it follows immediately that the temperature is a quantity that provides a status information about the "energy state" of the body.
We can refer to the zero law of thermodynamics: if there is thermal equilibrium, or there is no heat (Q) exchange between two bodies A and B placed in contact, the bodies A and B are at the same temperature !
This property is transitive: if there is a thermal equilibrium between the bodies A and B and between the bodies B and C, there is also thermal equilibrium between the bodies A and C, although not in contact …
However, the “operative methods” for the measurement of temperature are realized by means of the properties of bodies and materials and their change with temperature …
the variations of physical dimensions → the volume ΔV or the length Δl, according to the form of the body;
the variations of electrical properties → the resistance ΔR;
the variations of physical state → very useful to define the reference temperatures T .

## Slide 3
The gas thermometer experience:
applies
… at V = cost …
Given that we want Ts = Ti + 100 we write
so we get 0,36609·Ti = 100 and Ti = 273,16 [K]
= 1,36609 …
= 1,36609 …
Extrapolated,
For po=0
Gas removal

## Slide 4
For calibrating all industrial thermometers, six fixed temperature points are recommended, defined as the equilibrium temperatures between 2 phases between: solid/liquid/vapor at the standard pressure of 1 atmosphere (101325 Pa).
Oxygen point (-182.96°C): equilibrium T between O(L) and O(V)
Water triple point (+0.01°C): equilibrium T between H2O(S) , H2O(L) , H2O(V)
Water vapor point (+100°C): equilibrium T between H2O(L) e H2O(V)
Zync point (+419.58°): equilibrium T between Zn(S) e Zn(L)
Silver point (+961.93°C): equilibrium T between Ag(S) e Ag(L)
Gold point (+1064.43°C): equilibrium T between Au(S) e Au(L)
The official temperature scale (IS) is the thermodynamic scale where the temperature unit is the kelvin K but, it is also 1 K = 1 °C
ITS-90 Fixed-Point Cells
The International Temperature Scale of 1990 (ITS-90) is the “de facto” standard used in the industrialized world by international convention
Notes: Qualsiasi altra scelta romperebbe la relazione °K / °C. scala temperature termodinamiche ha base teorica ma la scelta effettiva è convenzionale

## Slide 5
Bimetallic Thermometers :
Mostly used as automatic thermal switches for safety temperature monitoring !

## Slide 6
Magnetothermic switch
5. bimetallic beam

## Slide 7
Electric industrial Thermometers :
Resistance temperature Detectors (RTD): Pt100
R=100  a 0 °C
Conductor
Reference thermometers
High Span, sensitivity, precision - Low…
RT

## Slide 8
Signal conditioning and correct circuit connections for RTD
Excludes resistor changes of the connecting wires
RTD
Rref=100
Rwire=1
Rwire=1
Rwire=1
white
white
red
red
+ sensed voltage
Exc current
Exc current
- sensed voltage
Rwire=1
Voltmeter
Current
source
I0
I0
Four-wire ohmmeter technique

## Slide 9
R4 is used to balance the bridge;
The adjustable resistance has no influence on the resistance of the bridge legs;
When using long wires subjected to temperature changes: minimizes the effects on the connecting wires due to the current (for low T)
When R1=R2=10R3
Signal conditioning and correct circuit connections for RTD
Notes: b) Più accuratezza, la variazione di resistenza non influeisce sulle gambe del ponte
d)

## Slide 10
3 wire circuit

## Slide 12
Thermistors
Large Negative Temperature Coefficient (NTC)
Highly non-linear: β=?
High sensitivity,
low precision: cheap
Semi-Conductor: made of polycrystalline mixed oxide ceramics, major material is Mn, Ni
Due to a “fair resistance changes” with temperature, they can be connected to a simple ohmmeter
Notes: Spesso B non è fornito

## Slide 13
made of polycrystalline ceramic on the basis of barium titanate – circuit protection
PTC – Positive Temperature Coefficient

## Slide 14
Thermocouples :
Pilot flame
Thermocouple
spring
Gas circuit

## Slide 15
Thermocouples :
if TB = TA
if TB > TA
Thermoelectric effects :

## Slide 16
SEEBEK effect: at the junctions of two isothermal metallic materials a potential difference is established. This is the main effect, and it is the one that originates the Δe when the two junctions are not at the same temperature.
PELTIER effect: if in the circuit of the thermocouple the junctions are at temperature TB > TA then with Δe ≠ 0, and the circulation of electric current is allowed, this tends to re-establish the thermal equilibrium, cooling the joint B at a higher temperature and heating the joint A at lower temperature.
THOMSON effect: if a conductor is not isothermal, a potential gradient appears on it.
ca

## Slide 18
Circuital connections:
Rule of the intermediate metals
It’s OK only if T3 = T4
It’s always OK !
Measure one T vs the other: Direct conversion of the fem on the table

## Slide 20
Thermocouple with electronic junction (with reference temperature Ta different from 0°C)
Rule of the intermediate temperatures
If Tref = Ta ≠ T0 = 0 °C the thermocouple detects Δe ∝ Tx − Ta that can not be interpreted directly on the tables !
The semiconductor thermometer measures Tref = Ta , the compensation circuit processes the Ecomp ∝ Ta − T0 and adds it to the Δe produced by the thermocouple.
The voltmeter receives the compensated voltage E = Δe + Ecomp that can be interpreted on the thermocouple tables.

## Slide 21
Electronic semiconductor temperature sensor:
Silicon transistor Q8 and Q11 produce the voltage
1 mm
1.6 mm
Notes: Silicon transistor voltage is temperature dependent: q8 and q11. R5 and R6 convert it to a current

## Slide 22
For all those situations in which the contact between the physical phenomenon and thermometer is NOT possible:
RADIATION thermometers
The thermal radiation of a body originates by thermal agitation of it’s atoms but, outside the body, it is nothing more than a regular electromagnetic radiation with a wavelength between 0,3 e 40 µm.
The spectrum of visible radiation is with a wavelength from 0.1 to less than 1 micron therefore, a large part of the thermal radiation lies in the infrared.
Every body above absolute zero in temperature emits radiation dependent on its temperature.
The bodies with an ideal thermal radiation are the blacks bodies, as they completely absorb the radiation hitting them and, at a given temperature, they emit the highest amount of heat radiation.
- for each frequency it emits the maximum amount of thermal radiative energy possible (not the same for each wl)
- It is a diffuse emitter: the energy is radiated isotropically, independent of direction.
Notes: MWIR: In guided missile technology the 3–5 µm portion of this band is the atmospheric window in which the homing heads of passive IR 'heat seeking' missiles are designed to work, homing on to the Infrared signature of the target aircraft, typically the jet engine exhaust plume. This region is also known as thermal infrared.
LWIR: The "thermal imaging" region, in which sensors can obtain a completely passive image of objects only slightly higher in temperature than room temperature - for example, the human body - based on thermal emissions only and requiring no illumination such as the sun, moon, or infrared illuminator. This region is also called the "thermal infrared".

## Slide 23
The physical law that describes this phenomenon is the Planck ‘s law :
which provides the spectral radiation level Lλ as a function of the wavelength λ of the radiation and of the absolute temperature T of the black body (into a hemisphere, per unit wavelength).
The figure shows a representation of the law for several values of T :
with increasing T the radiance Lλ (the area under the curve is the total emitted power) increases but shifts to smaller wavelengths λ …
The sensitivity?
Notes: L is the amount of radiation emitted from a flat surface into a hemisphere, per unit wavelength at the wavelength lambda
sr: steradianti. A steradian can be defined as the solid angle subtended at the center of a unit sphere by a unit area on its surface.

## Slide 24
The real bodies behavior differently from the black body and this fact is reckoned in terms of emissivity ε (dimensionless parameter):
where Ll is the radiance of the black body and Lla the radiance of the real body at the same temperature T of the black body. Real body radiance therefore can be written as :
Together with the emissivity e other two parameters are also considered: the reflectance r and transmittance q of surrounding bodies. If the body is in thermal equilibrium, the radiated energy is equal to the energy absorbed and we can write the following relationship :
Then we must also take into account the losses due to air or dust along the optical path, the object's size and its distance from the thermometer (area length error) …
a
×
If 0 : real body

## Slide 25
If the emitance of a material is known it can be compensated

## Slide 26
The detectors are divided into thermal detectors (bolometers) and photodetectors (pyrometers)
Fire sensors, intruder alarm… faster as the electrical output is based on the changes in detector temperature (pyroelectric effect – lithium tantalate crystals), not temperature itself
Absorb a maximum of the incoming radiation at all wl;
The absorbed radiation causes the detector (typically a thin metal layer - resistor) temperature to rise each 50x50µm pixel…
Electrical signal in response to the incoming radiant flux
Photon detectors
Notes: Pirometro : cristalli di litio talantato

## Slide 27
THERMOGRAPHY :
The cameras are all individually calibrated by comparison with a black body at a controlled temperature T .
For each pixel of the image, the voltage-temperature characteristic V-T is stored in the LUT memory
(Look-Up-Table)

## Slide 28
Thermal camcorder :