## Slide 1
Mechanical Measurements
Lesson 14

## Slide 2
Some VELOCITY measurement techniques …
linear and angular velocities
of a vehicle driving shaft …

## Slide 3
VELOCITY measurement in fluids
Section (1) pressure energy + kinetic energy
Section (2) total pressure energy (called stagnation)
The energy of the fluid is kept between sections (1) and (2) therefore, we can write the Bernoulli’s equation:
If in equation we substitute the Stevino’s law we get …
the first graduation curve: → which, however, is NOT linear !
Incompressible fluid: v << vs

## Slide 4
With a little improvement, this simple device can be employed also for gas flowing inside bigger pipes …
There are two ports that can be seen in the figure: the dynamic port Pd and the static port Ps
Note that here, the density of the fluid which flows in the pipe must be different (much smaller) than the density of the manometer fluid: the graduation curve is therefore …
Ps – presa statica
Pd – presa dinamica
Pd
Ps

## Slide 5
VELOCITY measurement for fluids flowing NOT in pipes
We have to understand where to correctly place the static and the dynamic ports …
Energy equations in sections S1 and S2 are :
S1)
S2) ,
Where is the generic pressure
where is the pressure coefficient
Kp is a crucial coefficient used to decide “where to put the ports” on the probe !

## Slide 6
Kp is maximum (= 1) when w = 0, or v = - vf and the speed of the fluid element is slowed down to zero. This situation occurs only on the tip of the probe, along the probe axis.
In that point the generic pressure p = Ptot is maximum. This is the right site to place the dynamic intake Pd .
Kp is null (= 0) when w = v, or vf = 0 where the surface of the tube does not cause any “interference" on the fluid element speed. In that point the generic pressure p = pst is the static pressure. This is the right site to place the static port Ps .
Note that the Kp can take on negative values in the tail regions of the probe, when w > v and p <pst , in the depression zones where the fluid accelerates to reform the undisturbed configuration of the fluid-dynamic field .
The graduation curve of such a probe is therefore:

## Slide 7
The real PITOT tube :

## Slide 8
if we use the Pitot tube coupled with a water differential manometer to measure the velocity of air, in the
graduation curve it is possible to measure directly the pressure just reading the mmH2O .
In fact, and
if we express the density of air in the technical system:

## Slide 9
Hot-wire anemometer
I
When V increases, h(V) increases and Tw and Rw decrease
A current I passes through a small tungsten wire: the wire warms up
The heat generated by the Joule effect is the same heat removed by thermal convection
If we work at Tw constant (Rw), and Tf is constant
Notes: Calore prodotto per effetto Joule=calore asportato per convezione

## Slide 10
0.001s lag
The probe is put in a fluid with a known V1;
The resistor RI is adjusted to set the hot-wire anemometer current I at a value low enough to prevent burnout, but high enough to give adequate sensitivity: Rw comes to a definite Tw and Rw;
R3 is adjusted to balance the bridge;
We get the first point of the calibration curve: ;
V changes to V2: Tw and Rw change unbalancing the bridge. The bridge is balanced changing RI (changing I): a new point is obtained
Notes: Calore prodotto per effetto Joule=calore asportato per convezione