## Slide 1
Mechanical Measurements
Lesson 15

## Slide 2
FLOW measurements are based on the principle of conservation of mass
No !
No !
Entering mass = outgoing mass (me = mu)
Mass flow:
Weight flow rate:
Volume flow rate:
Volume = S · Δx →
with velocity of the fluid “control” element !

## Slide 3
It’s known !
?
Laminar flow
Turbolent flow
In the case of incompressible fluid (ρ = constant), the speed distribution in the flow must be "measured" in some way and the average speed must be calculated some how; only then we can write:
Volumetric flow is always preferred when measuring flow in pipes :

## Slide 4
For flow speed pressure drop methods are employed
Volumetric flow Qv is constant in every section of the pipe
being S2 < S1 we get v2 > v1
VENTURI meter

## Slide 5
The BERNOULLI theorem is applied between the sections S1 and S2
Volumetric flow:
Weight flow:
The throttling coefficient is : and being

## Slide 6
The losses (always present) are taken into account by a discharge coefficient :
Actually, C = C(Re) is provided by the instrument producer ! (can be C>0.9)
If we do not have at our disposal the necessary length to insert in the pipe a Venturi flow meter, we can employ a diaphragm flow meter:
It has the same graduation curve of the Venturi meter but with a much smaller discharge coefficient C = C (Re) because of the considerable induced pressure losses (over 40%) happening at the diaphragm bottleneck !
Discharge coefficient:
Venturi meter are long to allow for a good recovery of pressure drop

## Slide 7
( NR = Re )
Diaphragm flow meter

## Slide 8
An intermediate solution can be the nozzle flow meter (with pressure losses around 20%):

## Slide 9
Rotameter flow meter :
The Graduation curve is based on the equilibrium of two forces:
1) weight (in water) of the floater:
2) thrust (push-up) of the flow:
with ρ* = ρc – ρf (floater density – fluid density)
Vc floater volume and A maximum floater cross section.
being all the quantities listed in the left term of the equation above "constant", the pressure difference between "below" and "above" the floater is also constant !

## Slide 10
If we apply the Bernoulli theorem we get:
flow velocity at the «annular section»
which is always constant !
If the flow velocity v is “constant” for each annular section, or for each height reached by the float, being when the flow Q changes there must necessarily be a “variation of the section S” ...
Therefore, we have to make the transparent tube with a “truncated cone shape” !
The graduation curve is : with S = At(h) – A

## Slide 11
Note in the figure that rotameter flow meter are often “coupled as a shunt” with bigger diaphragm flow meter …
You can "see" the flow;
They need the gravity force to work;
Specific for the liquid

## Slide 12
Laminar flow meter :
They are designed to "force" the flow to remain laminar!
They work for fluid media with Re <2000
They measure even very small flow rates from a few cc/h to several tens of m3/min
The graduation curve is the Hagen-Poiseuille law
They are often used to measure gaseous flow (medical) at low
pressure and non-stationary conditions !
PROS: linear scale between Q and Δp, large accurate range, insensitive to upstream and downstream flow interference, bi-directional flows ...
AGAINST: introduce considerable Δp, becomes clogged if the gas is not clean ...
Q: volume flow rate
D: capillary inside diameter (e.g. 0.1 mm)
µ: fluid viscosity
L: length between pressure taps (e.g. 120 mm)
p: force drop

## Slide 13
Electromechanical turbine flow meter:
In the hypothesis of “absence of friction and inertia”, in every infinitesimal time dt the fluid goes from (a) to (b) while the rotor goes from (c) to (b).
Therefore we can write the simple relationships:
Graduation curve:
The rotor just "get out of the way" of the fluid

## Slide 14
The K factor (pulse num./liter) for the ideal volumetric meter should be a constant.
In reality, especially for low flow rates, the rotation speed of the turbine is affected by the viscosity of the processed fluid (density and temperature).
There are "smart instruments" (with a temperature sensor and a microprocessor) that manage to “compensate” this change (span turndown), however, only with fluids for which the calibration of the instrument was carried out ...
Typical Specifications:
Span: 50 cc/min – 150 m3/min (liquid)
300 cc/min – 450 m3/min (gas)
high accuracy; non-linearity error = 0,1 %
Δp max = 20-70 kPa
1° order measuring system
(τ = 10 ms)

## Slide 15
Electromagnetic flow meters :
Based on the electromagnetic induction principle:
e = Blv
B → magnetic induction field [Wb/m2 ]
l → moving electrical conductor [m]
v → velocity of the electrical conductor [m/s]
In a conductive fluid which passes through a magnetic field, there is a “ion separation” in the flow tube that generates a distribution of electrical potentials. The highest values are at the ends of the flow tube diameter: V+ - V- = BDpv
In real instruments, the area of the magnetic field B application is limited and the electric potential values measurable at the diameter point are smaller (the fluid exposed to the magnetic field B tends to short-circuit the potential developed under the B field)
If a conductor of length l moves with a transverse velocity v across a magnetic field B, there will be forces that move the + charges toward one end and the – to the other

## Slide 16
If the extension of the B field is at least 3 diameters Dp , the effect of the lateral short-circuit is reduced at the center of the B field; there are located the two electrodes to measure the electric potential !
The flowmeter tube must be made of NON magnetic material in the measurement zone, so as not to distort the field lines B …. and must be made of NON conductive material so as not to short circuit the electrical signal e !

## Slide 17
The magnetic field B can be applied either:
alternating "ac" (50-60Hz: low polarization of the electrodes, low distortion of the velocity profiles, low signal shifts, stable amplification, but induction of an ac spurious signals on the measuring signal)
continuously switched "dc" (by square waves of a few Hz "auto-zeroing", slower than the ac instruments, with time constants τ = 6s vs 2s)
The limit of the fluid conductivity is: 0.1 mS/cm . Below, we don’t get any electrical signal anymore !
This method is used for the measurement of blood flow in the capillaries with diameters up to d = 1mm
They do not introduce any obstruction to the flow and are insensitive to fluid density variations, viscosity and flow disturbances ... as long as the profile of v remains symmetrical …
Additional problems may arise if the pipeline is not full, eg. for wastewater ...
The electric potential e = Blv causes a small current i between the electrodes, that passes through the conduction path of the fluid, which has a resistance R -> e = R · i
R can be estimated at 1/σd (if σ conductivity of the fluid = 200 µS/cm for water; d diameter of the tube = eg. 1 cm; then R = 5000 Ω) and must be measured during the calibration of the instrument.
R determines the load effect on the signal manipulation circuit connected downstream ...
Notes: S=Siemens = 1/Ohm

## Slide 18
Ultrasound flow meters :
The sound in a fluid (a longitudinal pressure wave with appropriate frequency) propagates with speed dependent on the type (E, ρ) and on the conditions (T) of the fluid itself !
If the fluid is in motion, the absolute velocity of the sound perturbation is the “algebraic sum” of its propagation velocity in the medium and that of the fluid itself.
An appropriate positioning of the sound emitter and receiver along the flow tube, may allow an accurate determination of the velocity v of the fluid and, therefore, of its volumetric flow rate Q = S · v
To obtain a tight and well defined acoustic field, ultrasonic wave packets are used with frequencies well above the audible (typically 10 MHz)

## Slide 19
Flight time
(a) - with still fluid:
for water c = 1520 m/s and if L = 30 cm
the transit time is t0 = 0,2 ms
- with fluid in motion with a velocity v :
therefore generally small (a fraction of μs) and with t0 not directly measurable
(b) with t1 in the same direction of the flow and t2 in the opposite direction, both measurable directly.
Double: higher accuracy on the measurement.
(c)
(d) with «clamp-on» transducers outside the pipe
Transducers that are employed: Almost only piezoelectric crystals (emitters/receivers)
The main measurement techniques employed are:
flight time
doppler
Notes: In b) ho un dt doppio, maggior accuratezza nella misura. Inoltre in a) dovrei avere t0

## Slide 20
Doppler (only for fluids with bubbles or with immersed particles):
The emitter sends an ultrasonic continuous wave (with frequency up to 10 MHz) that is reflected by the particles and measured by the receiver (we assume a uniform profile of v):
Nonetheless, the real measurement is more complex and is based on FFT post processing …
Pulsed Doppler systems are commonly used in modern medical ultrasound systems for non-invasive measurement of flow in blood vessels ...

## Slide 21
To measure the flow of fuels or of chemical medium the mass flow ratio is more suited ...
Qm measurement through the Coriolis effect
Qm measurement by heat transfer
t = k Qm